## Export Summary
This document was exported from Numbers. Each table was converted to an Excel worksheet. All other objects on each Numbers sheet were placed on separate worksheets. Please be aware that formula calculations may differ in Excel.
Numbers Sheet Name | Numbers Table Name | Excel Worksheet Name
Radiation Counts Log
| Table 1 | Radiation Counts Log - Table 1
| Statistics | Radiation Counts Log - Statisti
| Runs | Radiation Counts Log - Runs
| Runs-1 | Radiation Counts Log - Runs-1
| Runs-2 | Radiation Counts Log - Runs-2
| “All Drawings from the Sheet” | Radiation Counts Log - Drawings
Twilite Calibration
| Table 1 | Twilite Calibration - Table 1
| Runs | Twilite Calibration - Runs
| Runs-2 | Twilite Calibration - Runs-2
| Runs-2-1 | Twilite Calibration - Runs-2-1
| Runs-2-2 | Twilite Calibration - Runs-2-2
| Runs-2-1-1 | Twilite Calibration - Runs-2-1-
| Runs-2-1-1-1 | Twilite Calibration - Runs-2-11
| Runs-2-1-1-2 | Twilite Calibration - Runs-2-12
| “All Drawings from the Sheet” | Twilite Calibration - Drawings

## Radiation Counts Log - Table 1
DATE: | 2016-09-09 | PROJECT ID: | CCIR_00754
SUBJECT ID: | HYGLY28 | PRINCIPLE INVESTIGATOR: | Arbelaez
ISOTOPES: |  | DOSES DELIVERED / mCi: | 115.9
|  | OPERATOR: | JJL

## Radiation Counts Log - Statisti
ELAPSED TIME | MAX W-01 / KCPM | MAX CF / KDPM | MAX GE-68 / KDPM | Σ GE-68 / mCi | Σ WEIGHT SMPL / G
0.097141 | 408.3 | 599.1 | 4139 | 42.5599 | 29.6304
| MEDIAN W-01 / KCPM | MEDIAN CF / KDPM | MEDIAN GE-68 / KDPM
| 147.8 | 413.5 | 372.3
| MIN W-01 / KCPM | MIN CF / KDPM | MIN GE-68 / KDPM
| 29.16 | 86.22 | 0

## Radiation Counts Log - Runs
TUBE | Time / hh:mm:ss | COUNTS / cpm | counts S.E. / cpm | ENTERED | TRACER | TIME DRAWN / hh:mm:ss | TIME COUNTED / hh:mm:ss | W-01 / kcpm | CF / kdpm | Ge-68 / kdpm | MASS DRY / g | MASS WET / g | MASS SAMPLE / G | aperture corr Ge-68 / kdpm / g | TRUE DECAY & APERTURE CORR Ge-68 / kdpm / g | COMMENTS
26 | 41160.484 | 813 |  | FALSE | [18F]DG | 41160.5023 | 41160.5065 |  |  | 0 | 3.7736 | 4.3328 | 0.5592 | 0 | 0
27 | 41160.4854 | too high |  | FALSE | [18F]DG | 41160.5023 | 41160.5078 |  |  | 0 | 3.7716 | 4.4608 | 0.6892 | 0 | 0
28 | 41160.4979 | 304 |  | TRUE | [18F]DG | 41160.5024 | 41160.5097 |  |  | 0.111 | 3.8659 | 4.6719 | 0.806 | 0.13874 | 0.1508
29 | 41160.5049 | 235 |  | TRUE | [18F]DG | 41160.5025 | 41160.5151 |  |  | 3242 | 3.8124 | 4.5233 | 0.7109 | 4548.1318 | 5194.1394
30 | 41160.5403 | 245 |  | TRUE | [18F]DG | 41160.5025 | 41160.516 |  |  | 4139 | 3.8259 | 4.5253 | 0.6994 | 5895.6242 | 6789.239
31 | 41160.5723 | 223 |  | TRUE | [18F]DG | 41160.5026 | 41160.5169 |  |  | 2617 | 3.7517 | 4.5498 | 0.7981 | 3300.4262 | 3830.798
32 | 41160.6042 | 196 |  | TRUE | [18F]DG | 41160.5026 | 41160.5178 |  |  | 1682 | 3.7925 | 4.5422 | 0.7497 | 2246.2032 | 2628.6534
33 | 41160.5318 | 266 |  | TRUE | [18F]DG | 41160.5027 | 41160.5186 |  |  | 1555 | 3.8069 | 4.5149 | 0.708 | 2189.8089 | 2582.6926
34 |  |  |  | FALSE | [18F]DG | 41160.5027 | 41160.5195 |  |  | 1476 | 3.824 | 4.6065 | 0.7825 | 1895.1971 | 2253.41
35 |  |  |  | FALSE | [18F]DG | 41160.5028 | 41160.5223 |  |  | 1178 | 3.8202 | 4.6551 | 0.8349 | 1426.4257 | 1739.7857
36 |  |  |  | FALSE | [18F]DG | 41160.5028 | 41160.5232 |  |  | 960.2 | 3.7968 | 4.5934 | 0.7966 | 1213.0251 | 1491.229
37 |  |  |  | FALSE | [18F]DG | 41160.5029 | 41160.5382 | 408.3 |  | 929.9 | 3.7637 | 4.7233 | 0.9596 | 996.6374 | 1403.665 | error w-01 vs ge68
38 |  |  |  | FALSE | [18F]DG | 41160.5029 | 41160.5391 | 330.1 |  | 747.5 | 3.8539 | 4.6534 | 0.7995 | 941.2104 | 1337.0907 | error w-01 vs ge68
39 | 41160.5403 | 245 |  | TRUE | [18F]DG | 41160.503 | 41160.5415 | 293.9 |  | 663.5 | 3.9044 | 4.6843 | 0.7799 | 854.5318 | 1240.9515 | error w-01 vs ge68
40 |  |  |  | FALSE | [18F]DG | 41160.503 | 41160.5424 | 262.8 |  | 591.5 | 3.7845 | 4.5385 | 0.754 | 785.76 | 1150.6076 | error w-01 vs ge68
41 |  |  |  | FALSE | [18F]DG | 41160.5031 | 41160.5709 | 179.2 | 529.8 | 401.2 | 3.7657 | 4.4717 | 0.706 | 566.4785 | 1074.0592
42 | 41160.5723 | 223 |  | TRUE | [18F]DG | 41160.5031 | 41160.5733 | 181.5 | 536.7 | 405 | 3.7869 | 4.5556 | 0.7687 | 528.5531 | 1024.9776
43 |  |  |  | FALSE | [18F]DG | 41160.5032 | 41160.5744 | 187.5 | 554.3 | 420.2 | 3.8362 | 4.6699 | 0.8337 | 509.4717 | 997.482
44 |  |  |  | FALSE | [18F]DG | 41160.5033 | 41160.5754 | 172.5 | 509.9 | 388.6 | 3.7874 | 4.5812 | 0.7938 | 492.4933 | 973.2119
45 |  |  |  | FALSE | [18F]DG | 41160.5033 | 41160.5822 | 171.1 | 505.8 | 382.3 | 3.8288 | 4.6892 | 0.8604 | 450.6573 | 947.2897
46 |  |  |  | FALSE | [18F]DG | 41160.5034 | 41160.584 | 130.8 | 386.6 | 292.2 | 3.7812 | 4.4461 | 0.6649 | 436.4774 | 932.4742
47 |  |  |  | FALSE | [18F]DG | 41160.5034 | 41160.585 | 173.9 | 514 | 387.4 | 3.8163 | 4.7355 | 0.9192 | 430.8902 | 929.2006
48 |  |  |  | FALSE | [18F]DG | 41160.5035 | 41160.5861 | 162.8 | 481.4 | 362.3 | 3.8673 | 4.7304 | 0.8631 | 425.8946 | 927.6561
49 |  |  |  | FALSE | [18F]DG | 41160.5035 | 41160.5872 | 202.6 | 599.1 | 452.9 | 3.765 | 4.892 | 1.127 | 425.0892 | 935.2054
50 |  |  |  | FALSE | [18F]DG | 41160.5036 | 41160.5883 | 130.4 | 385.6 | 289.9 | 3.8394 | 4.5383 | 0.6989 | 413.2122 | 918.3067
51 |  |  |  | FALSE | [18F]DG | 41160.5036 | 41160.5894 | 163.1 | 482.9 | 361.9 | 3.8425 | 4.7426 | 0.9001 | 409.9685 | 919.8656
52 |  |  |  | FALSE | [18F]DG | 41160.5037 | 41160.5904 | 125.4 | 370.8 | 282.1 | 3.8599 | 4.5432 | 0.6833 | 410.6946 | 930.0686
53 |  |  |  | FALSE | [18F]DG | 41160.5037 | 41160.5925 | 139.9 | 413.5 | 311.6 | 3.7908 | 4.5838 | 0.793 | 395.2693 | 912.6387
54 |  |  |  | FALSE | [18F]DG | 41160.5038 | 41160.5938 | 123.9 | 366.2 | 276.6 | 3.8273 | 4.5285 | 0.7012 | 393.0454 | 917.8775
55 |  |  |  | FALSE | [18F]DG | 41160.5038 | 41160.5948 | 143.6 | 424.5 | 321.7 | 3.8292 | 4.6693 | 0.8401 | 387.3822 | 913.5499
56 |  |  |  | FALSE | [18F]DG | 41160.5038 | 41160.596 | 147.8 | 437 | 329.4 | 3.7886 | 4.6698 | 0.8812 | 380.1748 | 905.8476
57 |  |  |  | FALSE | [18F]DG | 41160.5069 | 41160.597 | 85.82 | 253.7 | 191.9 | 3.849 | 4.5799 | 0.7309 | 262.3573 | 631.2043
58 |  |  |  | FALSE | [18F]DG | 41160.509 | 41160.5981 | 69.21 | 204.6 | 154.55 | 3.8158 | 4.5303 | 0.7145 | 215.7967 | 524.2907
59 |  |  |  | FALSE | [18F]DG | 41160.5125 | 41160.5995 | 58.57 | 173.2 | 130.1 | 3.8624 | 4.5832 | 0.7208 | 180.1795 | 443.1336
60 |  |  |  | FALSE | [18F]DG | 41160.516 | 41160.6005 | 46.02 | 136 | 101.6 | 3.7516 | 4.4381 | 0.6865 | 147.2665 | 365.7496
61 |  |  |  | FALSE | [18F]DG | 41160.5229 | 41160.6016 | 39.78 | 117.6 | 88.69 | 3.787 | 4.523 | 0.736 | 120.4747 | 302.0572
62 |  |  |  | FALSE | [18F]DG | 41160.5368 | 41160.6026 | 30.16 | 89.15 | 66.51 | 3.747 | 4.5003 | 0.7533 | 88.4287 | 223.8207
63 |  |  |  | FALSE | [18F]DG | 41160.5437 | 41160.6036 | 29.16 | 86.22 | 64.91 | 3.7972 | 4.6235 | 0.8263 | 79.3328 | 202.5811

## Radiation Counts Log - Runs-1
TUBE | Time / hh:mm:ss | COUNTS / cpm | counts S.E. / cpm | ENTERED | TRACER | TIME DRAWN / hh:mm:ss | TIME COUNTED / hh:mm:ss | W-01 / kcpm | CF / kdpm | Ge-68 / kdpm | MASS DRY / g | MASS WET / g | Mass sample / g | aperture corr Ge-68 / kdpm / g | DECAY & APERTURE CORR Ge-68 / kdpm / g | COMMENTS
1 | 41160.3833 | 189 |  | TRUE | O[15O] | 41160.4343 | 41160.4375 |  |  | 442.8 | 3.7859 | 5.5693 | 1.7834 | 307.0506 | 1589.8942
P1 | 41160.4264 | 210 |  | TRUE | O[15O] | 41160.4343 | 41160.4407 |  |  | 21.07 | 3.8478 | 4.5595 | 0.7117 | 29.5277 | 748.0308 | Plasma
2 | 41160.4472 | 224 |  | TRUE | O[15O] | 41160.4764 | 41160.4807 |  |  | 126.3 | 3.8263 | 5.3021 | 1.4758 | 97.5833 | 734.6262
P2 |  |  |  | FALSE | O[15O] | 41160.4764 | 41160.4825 |  |  | 21.98 | 3.7965 | 4.4618 | 0.6653 | 32.8143 | 619.009 | Plasma

## Radiation Counts Log - Runs-2
TRACER | ADMINistration time / hh:mm:ss | True Admin. time / hh:mm:ss | dose / mCi | COMMENTS
C[15O] | 41160.4247 | 41160.4231 | 21
O[15O] | 41160.4357 | 41160.4341 | 18
H2[15O] | 41160.4466 | 41160.445 | 20.7
C[15O] | 41160.4583 | 41160.4567 | 15
O[15O] | 41160.4781 | 41160.4766 | 16
H2[15O] | 41160.4886 | 41160.487 | 20.3
[18F]DG | 41160.5021 | 41160.5005 | 4.9

## Radiation Counts Log - Drawings
RADIATION COUNTS LOG
BACKGROUND
SAMPLE
All times are local to facility and instrument. Decay corrections are to time of administration of tracer.
BACKGROUND
SAMPLE

## Twilite Calibration - Table 1
DATE: | 2016-09-09 | PROJECT ID: | CCIR_00754
CALC. DATE: | 2017-06-12 | PRINCIPLE INVESTIGATOR: | Arbelaez
|  | OPERATOR: | JJL

## Twilite Calibration - Runs
CLOCKS | Time Offset wrt NTS / +/-s
mMR console | 8.1019e-05
PMOD workstation | 0
mMR PEVCO lab | -0.0013657
CT radiation lab | 0
hand timers | 0.0015856
2nd PEVCO lab | 0

## Twilite Calibration - Runs-2
CYCLOTRON & DOSE CALIBRATOR | time / hh:mm:ss | dose / mCi | CYCLOTRON LOT ID | CYCLOTRON TIME | Cyclotron Activity / mCi / mL | Cyclotron volume / mL | Expected Dose / mCi | COMMENTS
syringe + cap dose | 41160.5233 | 3.02 | F1-090916 | 41160.2583 | 69.2 | 0.4 | 2.5197
residual dose | 41160.5233 | 0.5 | [5 empty] | guessing residual dose
net dose |  | 2.4889

## Twilite Calibration - Runs-2-1
PHANTOM | Original Volume / mL | Net Volume (phantom + dose) / mL | DECAY Corr Specific Activity / kBq/mL | COMMENTS
| 500 | 500 | 136.8108

## Twilite Calibration - Runs-2-2
WELL COUNTER
| Time / hh:mm:ss | COUNTS / cpm | counts S.E. / cpm | ENTERED | TRACER | TIME DRAWN / hh:mm:ss | TIME COUNTED / hh:mm:ss | W-01 / kcpm | CF / kdpm | Ge-68 / kdpm | MASS DRY / g | MASS WET / g | Mass sample / g | aperture corr Ge-68 / kdpm / g | DECAY & APERTURE CORR Ge-68 / kdpm / g | DECAY Corr Specific Activity / kBq/mL | S.A./S.A. OF DOSE CALIB. | COMMENTS
| 41160.5318 | 266 |  | TRUE | [18F]DG | 41160.5233 | 41160.53 |  |  | 2977 | 3.8113 | 5.8967 | 2.0854 | 1928.9765 | 1504.107 | 25.0684 | 0.18323 | failed mixing?
|  |  |  | FALSE | [18F]DG | [7 empty] | 0 | #DIV/0! |  | 0 | 0
|  |  |  | FALSE | [18F]DG | [7 empty] | 0 | #DIV/0! |  | 0 | 0

## Twilite Calibration - Runs-2-1-
TWILITE | Cath Place-ment time / hh:mm:ss | Enclosed Catheter length / cm | VISIBLE Volume / mL | Twilite baseline / coincident cps | Twilite loaded / coincident cps | Specific Count Rate / kcps/mL | Specific ACtivity/ kBq/mL | DECAY CORR specific activity / kBq/mL | S.A./S.A. OF DOSE CALIB. | COMMENTS
Medex REF 536035, 152.4 cm Ext. W/M/FLL Clamp APV = 1.1 mL | 41160.5573 | 20 | 0.14436 | 91.3 | 148.3 | 0.39485 | 116.6038 | 116.6038 | 0.8523
Medex | 41160.5573 | 20 | 0.14436 |  |  | 0 | 0 |  | 0
Medex | 41160.5573 | 20 | 0.14436 |  |  | 0 | 0 |  | 0

## Twilite Calibration - Runs-2-11
MMR | scan start time / hh:mm:ss | ROI mean / kBq/mL | ROI s.d. / kBq/mL | ROI Vol / cm3 | ROI PIXELS | ROI min / kBq/mL | ROI max / kBq/mL | DECAY Corr specific activity / kBq/mL | S.A./S.A. OF DOSE CALIB.
ROI1 | 41160.5574 | 103.61 | 22.81 | 518 | 129746 | 41.14 | 134.8 | 103.5337 | 0.75677
ROI2 | [7 empty] | 0 | 0
ROI3 | [7 empty] | 0 | 0

## Twilite Calibration - Runs-2-12
PMOD | Twilite specific activity / coincident kcps/mL | mMR SPECIFIC Activity / kBq /mL | REPORTED PMOD FACTOR | MAGIC PMOD FACTOR | COMMENTS
Calibration Window | 116.6038 | 103.61 | 1.996 | 147.95 | if well-mixed wrt mMR results, expect reported pmod factor = 0.429

## Twilite Calibration - Drawings
TWILITE CALIBRATION
BACKGROUND
SAMPLE